| INVITATION TO OFFER | Nr. | 8.428/09 |
| --- | --- | --- |

The bidder:

| Name: |  |
| --- | --- |
| Address: |  |

| Type of purchasing: |  | Service |  | Material | x | Equipment |  | Construction |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

We are kindly asking you to send us your Offer for:

| Nr. | Description of the goods |
| --- | --- |
| 1 | RAMAN SPECTROMETER as per enclosed technical description and specification |

The tender documentation is available on the internet:  http://www.ijs.si/ijsw/Objave.

Queries relating to the tender documentation can only be placed through the Public-Procurement Portal of the Official Gazette of the Republic of Slovenia: http://www.enarocanje.si

The bidders can submit their bids in Slovenian or English as per enclosed instructions for the bidders and tehnical specifications of the equipment. Variants are not permitted.

Due date for submission of the bids is 2 December 2009 by 11.00 a.m.

Address for submission of the bids: Institut »Jožef Stefan«, Jamova cesta 39, SI-1000  Ljubljana.

Tender documentation has to be submitted in one copy (tied together with a string and sealed in such a way that individual sheets or attachements cannot be added, taken away or replaced at a later stage) in an envelope, on which the following has to be clearly written:

- »NE ODPIRAJ – PONUDBA ZA RAMANSKI SPEKTROMETER« - the instruction in Slovene, according to which the envelope should not be opened by the mail service as it includes a bid for the raman spectrometer

- The exact address of the bidder

The public opening of the bids will be held on 2 December 2009 at 12.00 a.m. at the Jožef Stefan Institute, Jamova cesta 39, Ljubljana. The representatives wishing to be present at the opening, have to present a letter authorising them to attend the event to the selection board prior to the opening of the bids.

We look forward to receiving your offer.

| Ljubljana, | 28 October 2009 |  | Institut "Jožef Stefan" |
| --- | --- | --- | --- |
|  |  | Sales & Purchase Department, Head Jože Kašman, prof.mat. |

INSTRUCTIONS FOR THE BIDDERS

The public tender is carried out on the basis of Article 30 (point 1.a) of the Public Procurement Act (PPA-2) in a public call for tenders procedure with prior publication of a contract notice.

The bid has to prepared strictly using information as per enclosed specifications.

The bid can be submitted in Slovenian or English. The documents have to be put together in the following order:

- Offer  – Form 1

- Declaration on compliance with conditions – Form 2

- Statement from the bank – Form 3

- Sample contract (completed, signed and stamped sample contract, initialled on each page) – Form 4

The submitted forms have to be completed, signed and stamped. The entire documentation has to be printed and signed by the person authorised to represent the bidder.

The bid has to be valid until 31 January 2010.

When assessing the bids, the awarding authority’s selection board shall consider only complete bids, provided in compliance with tender specifications and all technical requirements. The most favourable proposal will be selected on the basis of the only criteria, the bidding price.

| Ljubljana, | 28 October 2009 |  | Institut "Jožef Stefan" |
| --- | --- | --- | --- |
|  |  | Sales & Purchase Department, Head Jože Kašman, prof.mat. |

Enclosure:  tehnical description and specifications of the equipment

TECHNICAL DESCRIPTION AND SPECIFICATIONS OF THE EQUIPMENT

Purpose: basic research and materials

Research field: inorganic chemistry, spectroscopy, materials

The subject of the public-tender is the supply of RAMAN SPECTROMETER with the following technical characteristics:

1.	SPECTROMETER

- Focal length: at least 800 mm

- Spectral resolution: better than 0.35 cm-1/pixel at 633 nm with a 1800 gr/mm grating

- Spectral range: at least 50 cm-1 to 4500 cm-1

- Gratings: at least two gratings 600gr/mm and 1800 gr/mm. Grating selection must be fully automated and software controlled

2.	DETECTOR

- CCD chip size: at least 25 mm in width and at least 1024x256 pixels. The full chip (1024 pixels) must be used to acquire the spectral data with no spectral distortion

- CCD chip:  Open Electrode CCD with quantum efficiency of at least 30% from 500nm to 800nm and a peak QE of at least 60%

- CCD cooling: by thermoelectric effect (Pelletier) down to -70°C

3.  FILTERS

- Edge filters: the Edge filters must operate in injection/rejection mode: the same Edge filter reflects the laser into the microscope and rejects the Rayleigh light

- Edge filters low-frequency limit: the filters must allow performing measurements down to 50 cm-1

4.	CONFOCAL MICROSCOPE

- Open microscope: The instrument must include an open microscope with free space under the objective turret, allowing  microscopic measurements on large objects

- Confocality: The instrument must be equipped with an automated true confocal aperture with continuously variable size

- Spatial resolution: The system must provide a lateral resolution of 1 µm or better and an axial (depth) resolution of 2 µm or better

- Coupling: the system must have a directly coupled microscope without fibber optics

- Integrated microscope: The microscope must be integrated in the main frame of the system to ensure highest spatial resolution, sensitivity and stability.

5.	MACRO SAMPLER

- Horizontal beam: The system should be equipped with a macro-sampler accessory with a horizontal beam, equipped with a lens of 40-mm focal length

- Large sample: The macro-sampler should allow to measure large samples. In particular, it must allow measuring Raman spectra of a sample at the bottom of a 30-cm glass or polymer tube positioned vertically, while the location of a sample of interest is 18 mm from the outer wall of the tube

- Liquid: The macro-sampler should be equipped with a cuvette holder and back reflecting mirror to obtain a multipass effect to enhance the Raman signal when measuring liquids

6. LASERS

- Laser lines: The system should be equipped with two laser lines: 532 nm and 633 nm (with high beam quality and circular beam profiles). The system must include all the necessary optics, and mechanical adaptations for coupling the lasers.

- Density filters: The system must allow to control laser beam power using software-controlled density filters

- Laser selection: The system must be fully automated, including laser switching. The system should not require periodic realignment of the lasers.

7. COMPUTER AND SOFTWARE

- PC: The system must include a dedicated PC and software for instrument control, data acquisition and handling.

- Scripts:  the software of the Raman system should allow the operator to program customized data acquisition or treatment routines using to control the different system functions, so one can define its own measurements routines according to the application.

- Software: Software for acquisition and handling of data should work in Windows XP Professional/Vista Ultimate environment. In case some other OS is supported, the connectivity with Windows OS should be provided, including the ability of data conversion into standard ASCII/binary type of files

- Upgrade: The policy of software upgrade should be defined

8. INSTALLATION, TRAINING OF OPERATORS AND SUPPORT

- Installation of equipment

- Training of operator(s)

- Free support over telephone

9.	OTHER REQUIREMENTS

- The system should be provided with tools kit including sample references, and full manuals.

- The system should be provided with one year warranty on parts and labour

- The system must be CE marked

Details about the bidder:                                                                                                FORM  1

| Name: |  |
| --- | --- |
| Address: |  |
| VAT identification number: |  |
| Registration number: |  |
| Bank account number and name of the bank: |  |  |  |
| Contact person: |  |
| Contact person's E-mail address: |  |
| Person responsible for signing the contract: podpis pogodbe: |  |

INSTITUT "JOŽEF STEFAN"

Jamova cesta 39

1000 Ljubljana

| OFFER Nr. Place: Date: |
| --- |
| On the basis of invitation to offer Nr. 8.428/09 of ...................... we are pleased to submitt the following offer: |
| Type of purchasing: |  | Service |  | Material | x | Equipment |  | Construction |

| No. | Description of the goods | Quantity | Price | Discount % | Value |
| --- | --- | --- | --- | --- | --- |
| 1 | RAMAN SPECTROMETER | 1 |  |  |  |

|  |  |  | TOTAL |  |
| --- | --- | --- | --- | --- |
|  |  |  | DISCOUNT |  |
|  |  |  | VALUE without VAT |  |

| Delivery date: |  |
| --- | --- |
| Payment terms (at least 30 days after issuing the Invoice): |  |
| Delivery term: | CIF Ljubljana (price includes delivery to the client's address) |
| Warranty period: |  |
| Validity of offer: | 31.1.2010 |

OBLIGATORY ENCLOSURE:  Catalogues with technical specifications

|  |  | The bidder: |
| --- | --- | --- |
|  |  | (signature of the representative) (name and surname) |

FORM  2

DECLARATION ON COMPLIANCE WITH CONDITIONS

The bidder:

| Name: |  |
| --- | --- |
| Address: |  |
| VAT identification number: |  | Registration Nr.: |  |
| Name of the Court/Register: |  |
| Number of entry in the register of companies: |  |

We hereby declare, that we comply with the following conditions:

| 1. | We accept all the requirements from the tender documentation. |
| --- | --- |
| 2. | We declare that our company neither our legal representative has never been convicted of the criminal offences determined in the Penal Code (the Official Gazette of the Republic of Slovenia, Nos. 55/08, 66/08). |
| 3. | We are not subject to enforced recovery procedure, bankruptcy or liquidation procedures, or any other similar procedure that would lead to the closure of our activities. |
| 4. | We have paid all the taxes and social-security contributions in line with the regulations of the country in which we are established. |
| 5. | We are licensed to provide the service that is the subject of this tender. |
| 6. | In the six months prior to the deadline for submitting the bid, our bank account(s) was (were) never frozen. |
| 7. | We have no outstanding liabilities to our subcontractors from previous public-procurement procedures. |
| 8. | Functional installation and testing start-up of the equipment will be provided by our qualified personnel. |
| 9. | We assure technical support by our qualified personnel, as well as on-site repair service that shall be carried out in a response time of 48 hours after the receipt of a notification. |

and are giving the following

C O N S E N T

The undersigned agrees that the Jožef Stefan Institute may, as part of the awarding procedure, obtain the data from the official records needed for checking the fulfilment of the requirements from the tender documentation relating to the RAMAN SPECTROMETER.

|  |  | The bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

FORM 3

# STATEMENT FROM THE BANK

Company:

Address:

As regards your letter informing us that you intend to submit your bid relating to the Jožef Stefan Institute’s public tender, with the code 8.428/09, regarding the supply of the RAMAN SPECTROMETER

S T A T E

that the company......................................................................, requesting the guarantees, will

obtain from us, without reservations, the following:

- the Performance Guarantee corresponding to 9% of the total contractual value without the VAT, valid for at least 30 days after the expiry of the deadline for the completion of the business deal;

- the Warranty Guarantee corresponding to 9% of the total contractual value without the VAT, valid for one additional day after the expiry of the warranty period set in the warranty certificates of the supplier.

The above guarantees can be obtained after presenting the relevant contract and in line with the business policy of the bank.

Place and date:

Bank stamp                                    Signature:

FORM 4

SAMPLE CONTRACT

Jožef Stefan Institute, Jamova cesta 39, 1000 Ljubljana, represented by the Director, Prof. Dr. Jadran Lenarčič, henceforth referred to as “the client”, and

.........................................................................., represented by the Director.............................

henceforth referred to as “the contractor”,

have agreed on the following

C O N T R A C T   No. …………..

Article 1

The contracting parties have established that, in line with the Public Procurement Act, the client awarded the contract to the most cost-effective bidder with regard to the realisation of the supply requirement expressed in the public tender published on the Public Procurement Portal of the Republic of Slovenia under the reference number JN…………………. .

With this contract the client and the contractor agree on their mutual rights and obligations, necessary for the smooth operation of purchasing and selling the following device:

RAMAN SPECTROMETER

specified, in detail, in the Offer …………....... as of ...………….  that is integral part of this contract.

Article 2

The value of the above equipment is …….………….. EUR. The price does not include the VAT.

Article 3

The delivery period for the equipment that is the subject of this contract is ………..... after the date of the contract.

The contractor shall provide free-of-charge transportation of the equipment and its functional installation.

The client shall sign the handover note after establishing that the correctly delivered equipment has been properly installed and that the contractor presented to the client the warranty certificates and the Warranty Guarantee.

Article 4

The contractor shall issue the invoices to the client in the following order:

- for 40% of the contractual amount after the signing of the contract,

- for 50% of the contractual amount after the contractor receives the equipment,

- for 10% of the contractual amount after the functional installation of the equipment, on the basis of the handover note signed by the client and the contractor.

The client is obliged to pay an invoice for the goods, i.e., the equipment, within………..days of the correct issuing of the invoice.

The contracting parties agree that, in the case of a late payment, the client has to pay the default interest.

Article 5

To guarantee a successful realisation of the contractual obligations, within 10 days of the conclusion of this contract, the contractor shall submit to the client the Performance Guarantee corresponding to 9% of the total contractual value, without the VAT, which has to be valid for at least 30 days after the functional installation of the equipment.

With respect to correcting any faults within the warranty period, when signing the handover note, the contractor shall submit to the client the Warranty Guarantee corresponding to 9% of the total contractual value, without the VAT, which has to be valid for one additional day after the expiry of the warranty period determined in the warranty certificate.

Article 6

The responsibility for keeping the equipment undamaged is transferred from the contractor to the client once the equipment is delivered to the client.

The contractor provides the client with a warrantee to cover the quality of the equipment and its smooth operation at least to an extent equal to the extent of the warrantee given by the producer of the equipment to the contractor, i.e., ……years following the date of the receipt of the equipment.

During the warrantee period the contractor shall provide for an on-site repair service that shall be carried out in a response time of 48 hours after the receipt of a notification.

If the equipment has not been repaired within the above period, the contractor shall provide, for the time during which the equipment is being repaired, an equivalent substitute piece of equipment or a part of the equipment. In this case the warranty period shall be extended by a time equal to the duration of the repair work.

All transport costs and other costs relating to the repair work made during the warranty period are charged to the contractor.

Article 7

The contractor declares that neither the contractor nor anybody else acting on the contractor’s behalf did promise, offer or give any reward to an employee of the client or another person that might have influenced the client’s decision during the procedure of awarding the public contract with the purpose of influencing the content of the contract, any act or decision of the client relating to the tender before, during or after the selection of the bidder.

For the cases described in the previous paragraph, an already concluded and valid contract becomes void; and if the contract has not yet been validated, it is considered not to be concluded.

Article 8

The contracting parties shall determine all the amendments to this contract, after a previously reached agreement, with the annexes to this contract.

The contracting parties shall resolve any disagreement relating to the realisation of this contract by common consent.

In the case that the contracting parties cannot settle a dispute by common consent, they shall abide by the decision of the competent court in Ljubljana.

This contract comes into force on the day when it is signed by both contracting parties, provided that the contractor submits the Performance Guarantee within the deadline determined in Article 5 of this contract.

The contract is drawn up in four identical copies so that each contracting party receives two copies.

| CLIENT: |  | CONTRACTOR: |
| --- | --- | --- |
| Jožef Stefan Institute |  |  |
| Director: |  | Director: |
| Prof. Dr. Jadran Lenarčič |  |  |
| Date: |  | Date: |