PUBLIC-TENDER DOCUMENTATION

IN ACCORDANCE WITH THE LOW-VALUE PROCEDURE

FOR

»SUPPLY OF A REFLECTRON STAGE

OF TOF MASS SPECTROMETER FOR MeV-SIMS«

Internal reference number of the public tender: JN7/2019

Ljubljana, April 2019

CONTENTS

- Invitations to tender

- Instructions for the bidders

- Technical documentation of the public tender

- The bid (Proforma invoice) – FORM 1

- Details about the bidder – FORM 2

General data about the bidder and the consortium – FORM 2.1

Acting with or without subcontractors – FORM 3

The subcontractor’s authorisation relating to direct payments made by the contracting authority to the subcontractor(s) and consent – FORM 3.1

- Forms for establishing the capacities of a bidder

- FORM 4.1 – Declaration of compliance with the conditions for the economic

operator

- FORM 4.2 – Statement of the bidder for obtaining personal information from

official records

- FORM 4.3 – Declaration  of compliance with the conditions for the pursuit of a

professional activity

- FORM 4.4 – Statement of technical suitability

- FORM 4.5 – List of references

- FORM 4.6 – Statement of the reference client

- Sample contract – FORM 5

- INVITATION TO TENDER

The contracting authority, the Jožef Stefan Institute, Jamova cesta 39, 1000 Ljubljana, tax number SI55560822, registration number 5051606, bank account: 01100-6030344242 registered with the Public Payment Administration (PPA) Ljubljana, invites interested bidders to submit their written tender in accordance with this public tender documentation.

Subject of the contract is a »SUPPLY OF A REFLECTRON STAGE OF TOF MASS SPECTROMETER FOR MeV-SIMS«.

The public tender is carried out in accordance with Article 47 of the Public Procurement Act (official consolidated text PPA-3, the Official Gazette of the Republic of Slovenia, Nos. 91/2015 and 14/2018 hereinafter PPA-3) as a low-value public procurement procedure.

We invite you to submit your bid for this tender in accordance with the instructions for bidding.

The bids have to be prepared in line with the tender documentation, fulfilling all the conditions for the participation in this public tender.

| Ljubljana, April 2019; Institut “Jožef Stefan” Ljubljana. |  |  |
| --- | --- | --- |

- INSTRUCTIONS FOR THE BIDDERS

  - THE CONTRACTING AUTHORITY

The jožef stefan institute is inviting all interested bidders to submit their bids in line with the tender documentation.

  - the code and subject of the public tender

Internal reference No.: JN7/2019

Subject: »SUPPLY OF REFLECTRON STAGE OF TOF MASS SPECTROMETER

FOR MeV-SIMS«

  - implementation of the tender

The public tender is carried out on the basis of Article 47 of PPA-3 (Official Gazette of RS, no. 91/2015 and 14/2018) as a low-value public procurement procedure. The contracting authority shall select the bid and make a contract with the respective bidder on the basis of the conditions and criteria described below.

Before the deadline for the submission of tenders, the contracting authority may amend the tender documents. The contracting authority shall make all the amendments to the tender documentation 5 days before the current deadline at the latest. Any such amendment shall be an integral part of the tender documentation and shall be submitted through the Public-Procurement Portal. If necessary, the contracting authority shall extend the deadline for submitting the bids to make it possible for the bidders to consider the amendments.  In the case of an extended deadline for submitting the bids all the rights and obligations of the contracting authority and the bidders shall relate to the new dates arising from the extended deadline.

  - ACCESS to the tender documentation

The forms from the tender documentation (TD) are available at http://www.ijs.si/ijsw/Objave.

  - the bidding procedure

Any individual or legal entity registered for the service that is the subject of this tender can submit a bid to this tender.

2.6 THE LANGUAGE OF THE BIDS

The bidders should submit their bids in Slovene or English.

  - PREPARATION OF THE BIDS

To prepare their bids, the bidders should give all the details required with the attached specification. The bids should be valid till 31 July 2019.

2.8  A JOINT BID

A bid may be submitted by a group of contractual partners. This group shall submit a Partnership Agreement stipulating how the public contract awarded within this public-procurement procedure will be executed, specifying, in particular, the leading partner authorized to assume, from the Contracting Authority, the obligations, instructions, and possibly also the payments on behalf of, and for the account of, all the partners, as well as the share and type of services/goods to be provided by each individual partner. The Agreement should clearly stipulate that, in relation to the Contracting Authority, all the partners are collectively and individually liable for the realisation of the entire commitment and each part of it, in accordance with paragraph 3 of Article 81 of the PPA-3.

All the partners in a group have to meet the conditions for the recognition of their competencies with respect to their legal statuses (they must meet all the conditions in section 2.11) and the first point requirement 2.12.1. PROFESSIONAL CAPACITY OF THE BIDDER the compliance with the contracting authority's other conditions for the recognition of competencies shall be examined collectively for the entire group of partners.

2.9  A BID INVOLVING SUBCONTRACTORS:

In a subcontracting relationship the main contractor transfers the execution of the whole or a part of the awarded public contract to a third party, namely, a subcontractor. The definition of a subcontractor is given in point 1 of Article 94 of the PPA-3. The bidder could perform the award of the contract alone or together with subcontractor(s).

If the tenderer offers the execution of the contract by subcontractor(s), it is required to:

- provide the list of all subcontractors (each subcontractor should meet all the requirements under point  2.11) and the type of work that the subcontractor will be undertaking,

- provide the list of contact information and the legal representatives of the subcontractors,

- complete ESPD forms for each subcontractor,

- complete FORM 2.1 – The subcontractor’s authorisation to direct payments, if subcontractor required.

The main contractor will be required during the realization of the public tender to inform about any changes to the existing information specified in the preceding paragraph and to send information about new subcontractors, that are planned to be subsequently integrated into carrying out the work no later than 5 days after the change. In the case of the involvement of new subcontractors the main contractor must together with the notice provide the information and documents, specified in the second, third and fourth indents of the preceding paragraph and meet all the requirements under point 2.11.

The Contracting Authority will reject any subcontractor, if there are grounds for exclusion from the first, second or fourth paragraphs of Article 75 of the PPA-3, except in the case from the third paragraph of Article 75 of the PPA-3. The Contracting Authority may reject any subcontractor, if there are grounds for exclusion from the sixth paragraph of Article 75 of the PPA-3. The Contracting Authority will not accept the proposal to replace the subcontractor or the inclusion of a new subcontractor, if this has an impact on the smooth implementation or completion of the works and if the new subcontractor does not meet the conditions by the contracting authority in the tender documentation. In the case of rejection of the new subcontractor the Contracting Authority will inform the main contractor no later than 10 days from receipt of the proposal.

In the event that the subcontractor in accordance with and in such a way specified in the second and third paragraph of Article 94 of the PPA-3 requires direct payment, a direct payment to the subcontractor shall be deemed to be obligatory in accordance with this Act and bound by an obligation of the Contracting Authority and the main contractor.

Where the tenderer intends to carry out the contract with the subcontractor, which requires a direct payment in accordance with this Article:

- the main contractor in the contract shall authorize the contracting authority, on the basis of an approved invoice or situation by the main contractor to pay directly to the subcontractor,

- the subcontractor submits the consent on the basis of which the Contracting Authority instead of the tenderer settles the subcontractor's claim against the tenderer,

- the main contractor to its invoice or situation attaches an invoice or situation of a subcontractor, which was previously approved.

If direct payment to the subcontractor is not required, the contracting authority of the main contractor will require that not later than 60 days from the payment of the final invoice or situation the main contractor sends a written statement and a written statement from the subcontractor that the subcontractor has received payment for supplied goods directly related to the subject of the contract.

If the main contractor does not comply with Article 94 of PPA-3, the contracting authority to the National Review Commission submit a proposal on the introduction of the offence’s procedure from point 2 of the first paragraph of Article 112 of the PPA-3.

In the event that the tenderer acts with the subcontractors, the tenderer shall submit the authorisation to direct payments (FORM 3.1) or a statement that the bid does NOT act with subcontractors. In this case the tenderer submits a declaration of the bid without subcontractors (FORM 3).

2.10  COMPLETION OF THE BID – ADMISSIBLE TENDER:

»Admissible tender« shall mean a tender that is submitted by a tenderer or there are no grounds for exclusion and which meets the selection criteria, which meets the needs and requirements of the contracting authority set out in the technical specifications and the procurement documents, which was received in due time, regarding which there is no evidence of collusion or corruption, which has not been found by the contracting authority to be abnormally low, and whose price does not exceed the contracting authority’s budget.

Tender documentation must consist of the forms submitted in the following order:

  - FORM 1 – The bid (Proforma invoice)

  - FORM 2 – Details about the bidder

  - FORM 3 – Acting with subcontractors

  - forms for establishing the capacities of a bidder – FORMS 4.1 to 4.6 together with the attached proofs of fulfilling the conditions

  - the completed, signed and stamped sample contract initialled on each page – FORM 5

Only when acting as consortium and/or with subcontractors the bidder has to submit the following forms:

  - FORM 1 – The bid (Proforma invoice)

  - FORM 2.1 – General data about the bidder and the consortium

  - Partnership Agreement (only in the case the bidder acts as a party in a joint venture or a consortium of the bidders)

  - FORM 3 – Acting with or without subcontractors

  - FORM 3.1 – The subcontractor’s authorisation to direct payments made by the contracting authority to the subcontractor(s) and consent

  - forms for establishing the capacities of a bidder – FORMS 4.1 to 4.6 together with the attached proofs of fulfilling the conditions

  - FORM 5 – the completed, signed and stamped sample contract initialled on each page

In line with point 6 of Article 91 of PPA-3, the selected bidder shall submit, within 8 days of the receipt of the request, the details about:

- its founders, partners, including silent partners, shareholders, limited partnerships or other owners, as well as the equity shares of these entities;

- business entities that are considered to be associated with the selected bidder on the basis of the law governing companies.

The tenderer submitting the offer accepts criminal and material responsibility that all data and documents included in the tender are true. Otherwise, the tenderer shall be liable to the Contracting Authority for any damage caused to it.

2.11  GROUNDS FOR EXCLUSION OF THE BID

| 1. | The economic operator must be excluded by the contracting authority from participation in a public procurement procedure in the case it is found out, where the economic operator itself or any person who is a member of its administrative, management or supervisory body or has powers of representation, decision or control therein has been the subject of a conviction by final judgment which has the elements of the criminal offenses that are defined in the first paragraph of Article 75 of the PPA-3 In the case that the economic operator is in the situation referred to in the above paragraph, in accordance with paragraph 9 of Article 75 of the ZJN-3, the contracting authority may submit evidence that he has taken sufficient measures to prove his reliability despite the existence of reasons for exclusion. The proof: FORM 4.1 – Completed Declaration of compliance with the conditions for the economic operator for all economic operators in the offer (also for potential subcontractors and/or partners from a joint bid) Prior to signing the contract the contracting authority shall ask the selected bidder to submit the authorization to obtain the information from the criminal record (for the economic operator and for all persons who are members of the administrative, managerial or supervisory body of the economic operator or who the power to represent or to decide or control it). The bidder himself can attach the certificates from the criminal record. The certificates thus submitted must be legally relevant with respect to the conditions on the day of submitting the bid. |
| --- | --- |
| 2. | The economic operator must be excluded by the contracting authority from participation in a public procurement procedure in the case it does not fulfil the obligations related to taxes and other monetary non-fiscal obligations pursuant to the law regulating of the financial administration, collected by the tax authority in line with the provisions of the country in which we have our registered office or in Member State of the Contracting Authority and we have no outstanding liabilities as at the day of submitting the bid worth EUR 50 or more. On the day of submitting the bid, the economic operator must submit all accounts of tax deductions for the incomes arising from employment relationships for the period of the last five years. The proof: FORM 4.1 – Completed Declaration of compliance with the conditions for the economic operator for all economic operators in the offer (also for potential subcontractors and/or partners from a joint bid) |
| 3. | The economic operator must be excluded by the contracting authority from participation in a public procurement procedure in the case it is on the deadline for the submission of bids excluded from the public award procedures due to being included in the record of entities with negative references, as defined in point a) of the fourth paragraph of Article 75 of the PPA-3 The proof: FORM 4.1 – Completed Declaration of compliance with the conditions for the economic operator for all economic operators in the offer (also for potential subcontractors and/or partners from a joint bid) |
| 4. | The economic operator must be excluded by the contracting authority from participation in a public procurement procedure in the case it has been in the last 3 years prior to the expiry of the deadline for submitting the bids fined twice for an offence related to the payment for work, about working hours, on rest, on the performance of work on the basis of civil law contracts, despite the existence of elements of an employment relationship or in connection with the employment of illegal workers by way of a final decision of the competent authority of the Republic of Slovenia or another Member State or a third country. The proof: FORM 4.1 – Completed Declaration of compliance with the conditions for the economic operator for all economic operators in the offer (also for potential subcontractors and/or partners from a joint bid) |

The Contracting authority shall in accordance with the eight paragraph of Article 75 of the PPA-3 at any time during the procedure exclude an economic operator where it transpires that, in view of acts committed or omitted, the latter was or is either before or during the procedure in one of the situations referred to this instruction.

2.12   CONDITIONS FOR PARTICIPATION

2.12.1. PROFESSIONAL CAPACITY OF THE BIDDER

| 1. | The bidder is for the occupation that it takes in the tender registered in one of the professional or trade registers, that are kept in the Member State in which the economic operator is established. The list of professional or business registers in the Member States of the European Union is published in Annex XI of Directive 2014/24 / EU. The proof: FORM 4.3 – Completed Declaration of compliance with the conditions for the pursuit of a professional activity The Contracting authority reserves the right to verify the existence and content of the tender in case of doubt about the validity of the tenderer's statements. For this purpose, the declaration must contain all the necessary information in order for the Contracting authority to verify compliance with the condition in question in the official records. In the event that such a verification is not possible, the Contracting authority shall require the economic operator to submit a copy of the entry in one of the professional or business registers. |
| --- | --- |

2.12.2. TECHNICAL CAPACITY OF THE BIDDER

| 1. | The bidder guarantees the required technical suitability of the equipment in line with the technical requirements from the description of the equipment included in the tender documentation. The proof: FORM 4.4 – Completed Statement of technical suitability |
| --- | --- |
| 2. | The bidder has is the past 5 years (2014, 2015, 2016, 2017, 2018) sold (which means qualitative and quantitative in according with the contract) at least 3 TOF mass spectrometers with reflectron stages. The proof: FORM 4.5 – Completed »List of references« (FORM 4.3) FORM 4.6 – Completed confirmation of the business partners (reference clients) on the original form |

The contracting authority shall obtain the data from the official records by itself and on the basis of the received consent of the bidder. The contracting authority can also consider the data from the official records that the bidder obtained or submitted during the other public-tender procedures if these records are not older than 4 months.

If the country where the bidder is established cannot issue the proofs, or documents, required by the tender documentation, the bidder can, in line with Article 42 of PPA-3, submit an appropriate sworn statement given by the bidder or by a witness in the presence of a judicial or administrative authority, a notary, or an authorised body representing the commercial entities in the bidder’s country.

2.13 THE BIDDING PRICE

All the prices have to cover all the costs for the realisation of the relevant public contract. The final price has to include all its constituent elements. The VAT has to be indicated separately (for the bidders from Slovenia).

2.14  EXAMINATION AND EVALUATION OF THE BIDS:

Before awarding the contract, the Contracting authority will verify the existence and content of data or other information indicated in the tender of the tenderer to which he has decided to award the contract. The Contracting Authority will perform examination and evaluation in line with Article 89 of the PPA-3.

2.15  CRITERIA FOR ASSESSING AND EVALUATING THE BIDS:

The contracting authority will award the contract based on the most economically advantageous tender. Timely submitted bids will be assessed in accordance with the criteria specified below.

| No. | Criterion |  | No. of points |
| --- | --- | --- | --- |
| 1 | The bidding price: A = average price of all the bids that meet the requirements. The prices are assessed on the basis of their deviation from A | A ± 3% | 15 |
|  |  | more than 3% to 6% | 10 |
|  |  | more than 6% to 9% | 5 |
|  |  | more than 9% | 0 |
|  |  | less than -3% to -6% | 20 |
|  |  | less than -6% to -9% | 25 |
|  |  | less than -9% | 30 |
| 2 | Delivery time in months | 6 months or below | 5 |
|  |  | 7 do 12 mesecev | 3 |
|  |  | nad 12 mesecev | 0 |
| 3 | Striking energy of secondary ions at the sensitive surface of the ion detector at extraction voltage of 3 keV at the sample | For each additional 1 keV of energy above 11 keV additional 10 points, up to maximum 40 points. | max 40 |
| 4 | Set-up lighter than 40 kg | weight 40 kg and more | 0 |
|  |  | Weight below 40 kg | 10 |
| 5 | The number of references from the past five years relating to the supply of mass spectrometers TOF with reflectron stage | more than 3 references provided: | 10 |

The bidder obtaining the most points out of 95 possible points shall be asked to realise its bid.

- example of bids with equal to the best results:

In the event that two or more bidders after the evaluation received the same number of points, the contracting authority will select the one which will offer the configuration with higher secondary ion striking energy at the sensitive surface of the detector at the end of the flight path.

If they are after that two offers tied, the offer with a lower contract price will be selected.

2.16  QUERIES RELATING TO THE TENDER DOCUMENTATION

Communication with the bidders concerning the queries about the tender and the preparation of a bid shall be organised through the Public-Procurement Portal of the Official Gazette of the Republic of Slovenia: www.enarocanje.si.

2.17 DEADLINES FOR ANSWERING THE QUERIES ABOUT THE TENDER DOCUMENTATION

Queries about the tender documentation or any other query about the call for tender placed through the Public-Procurement Portal by 16:00 on 30 April 2019 shall be considered as timely. After this time, the contracting authority shall nor answer any queries about the tender.

2.18  SUBMISSION AND DELIVERY OF THE BID:

Bidders must submit their bids through the IT ePonudbe.si system at the web address http://eponudbe.si/en

Before submitting their bids, Bidders must register at the web address http://eponudbe.si/. If the Bidder is already registered in the IT ePonudbe.si, they must sign in at the same address.

A user, that is in the ePonudbe.si information system and authorised to submit tenders for an economic operator, submits a tender by clicking the »Submit tender« button. The ePonudbe.si information system records the identity and the time of the submission of the tender. The user by submitting the tender expresses and states his will to submit a binding tender in the name of the economic operator (Article 18 of the Obligations Code, Official Gazette of the Republic of Slovenia, no. 97/07 – uradno prečiščeno besedilo, 64/16 – odl. US and 20/18 – OROZ631). With its submission, the tender is binding for the time stated in the tender except if the user withdraws or changes it before the time for submitting the tenders has expired.

A bid is deemed to be submitted on time if the Contracting Authority receives it via the ePonudbe.si system http://eponudbe.si/en  by 7 May 2019 at 10.00 AM. A bid is deemed submitted if it is marked with the status "SUBMITTED" (“ODDANO”) in the IT ePonudbe.si system.

Bidders may withdraw or change their bids by the deadline for submission of the bids. Withdrawal of the bid can be performed on the ePonudbe.si system in the profile of the bidder that has submitted the bid to the system, when the bid is submitted, with the optional statement about the reason for the withdrawal, the change of the bid is performed within the same system first by withdrawal of already submitted bid, and afterwards with submission of a new bid before expiry of the deadline.

It is not possible to submit any bid after the expiry of the deadline for the submission of bids.

The page for submitting an e-bid in this e-procurement procedure can be accessed here:

https://eponudbe.si/en/Public-contracts/Current/Individual-tender/t/Nakup-reflektronske-stopnje-masnega-spektrometra-TOF-za-MeV-SIMS-Supply-of-a-reflectron-stage-of-tof-mass-spectrometer-for-MeV-SIMS/n/8332d2eabea4b622e44dcc2dcd3e04d5

or http://eponudbe.si/en

2.19 THE PUBLIC OPENING OF THE BIDS

Bids shall be opened automatically in the IT ePonudbe.si system on  7 May 2019 and will start at 10.01 AM at the web address http://eponudbe.si/en.

The opening takes place in the following way: at the hour set for the public opening of bids, the IT ePonudbe.si system automatically displays data on the Bidder and variants, if requested or allowed, and enables access to the PDF file that the Bidder uploaded in the ePonudbe.si system under the »Proforma Invoice« (»Predračun«) section. The public display automatically ends after 60 minutes. The Bidders that submitted bids have these data available in the IT ePonudbe.si system under the "Minutes on the bid opening" (»Zapisnik o odpiranju ponudb«) section.

- 2.21  form »bid (Proforma invoice)«

The Bidder must offer all listed items from the Form 1 – Bid (Proforma invoice), considering the technical description and the specification of the equipment, which is part of the Tender Documentation.

The Bidder completes all the listed items in the Proforma invoice, to a maximum of two decimal places.

If the Bidder declares a price of zero (0), it is understood this item is offered free of charge.

The Bidder must not change the content of Form 1.

The price must include all discounts and costs.

Where the Contracting authority, during the examination and evaluation of tenders, finds that a calculation error has occurred, it may correct it in line with Article 89, paragraph 7, of the PPA-3.

The Bidder uploads in the IT ePonudbe.si system under the section »Pro-forma Invoice« (»Predračun«) in a .pdf file the Proforma Invoice.

Bidder must also submit other documents (requested in point 2.10 of this Instruction for the bidders).

Bidder must scan all the bid documents and submit them in the PDF form in their e-bid.  Before that, the forms, where requested, must be completed, signed and stamped.

2.21 VALIDITY OF THE BID

The bid has to be valid untill 31 July 2019.

Variants are not permitted.

2.22 SIGNING OF THE CONTRACT

In line with the 6th paragraph of Article 14 of the Integrity and Prevention of Corruption Act (the Official Gazette of the Republic of Slovenia, Nos. 45/10, 26/11 and 43/11; ZIntPK), the selected bidder is obliged to submit, prior to signing the contract and at the request of the Jožef Stefan Institute, the awarding authority, a statement or details about the natural or legal entities owned by the selected bidder, including the silent partners and business entities that are considered to be associated with the selected bidder in line with the provisions of the law governing companies. If the bidder submits a false statement or gives untrue information about the required details, the contract shall be annulled.

The selected bidder shall sign the contract within 8 working days of the receipt of the contract signed by the contracting authority.

Prior to signing, the contract shall be modified depending on whether the selected bidder submits a joint bid, a bid with subcontractors, a bid consisting of different realisation phases and the like.

2.23 legal instruction:

A review claim referring to the contents of the call and/or the tender documentation, may be filed in 10 working days from the date of publication of the call for tender or the notice of supplementary information, information on the pending procedure or correction if this is a notice to amend or supplement requirements or criteria for selecting the most advantageous tenderer from the tender documentation or previously published contract notice; the review claim may refer to the amended, supplemented or clarified content of the publication or the tender documentation or a directly related reference to it in the original publication or tender documentation. A review claim may not be submitted after the deadline for submission of the bid, unless the deadline submission of the bid is less than 10 working days. In this case, a review claim may be filed within 10 working days of the date of publication of the contract notice.

The applicant shall pay a tax of 2.000,00 euros to the bank account of the Ministry of Finance, No. SI56 0110 0100 0358 802, opened at the Bank of Slovenia, Slovenska 35, 1505 Ljubljana, Slovenia, SWIFT CODE: BS LJ SI 2X; IBAN: SI56011001000358802 – a tax for a revision of public procurement.

A review claim shall be filed in writing directly to INSTITUT »JOŽEF STEFAN«., Jamova cesta 39, 1000 Ljubljana or sent by post as registered mail with an acknowledgement of receipt.

Ljubljana, April 2019 		    							    					           	Institut  »Jožef Stefan«

Ljubljana

3. TECHNICAL DESCRIPTION AND SPECIFICATIONS OF THE EQUIPMENT

The subject of public tender is supply of the Reflectron stage of the TOF Mass Spectrometer for MeV-SIMS spektroscopy, which will be used for molecular mapping in biological materials with high-energy focused heavy ion beam, i.e. 6 MeV Cl ions, at the measuring station of the high energy focused ion beam at Jožef Stefan Institute. The Reflectron stage does not require any mechanical support, as it will be installed as an extension stage at the existing linear Time-Of-Flight mass spectrometer. The reflectron stage must be equipped with a secondary detector with postacceleration.

TECHNICAL CHARACTERISTICS:

- Reflectron stage of the TOF Mass Spectrometer (Short: Reflectron stage) should enable incremental upgrade of the existing linear Time-Of-Flight (TOF) mass spectrometer for MeV-SIMS at Jožef Stefan Institute, using existing retracting/supporting mechanics upon the simple adaptation.

- Reflectron stage should comprise of the appropriate drift space with a selectable potential (Field Free Liner),  a device with reflecting electrostatic field, and detector of secondary ions at the end of the secondary ion path. The Reflectron stage is a part of high vacuum system, and its components vacuum leak-proof.

- The detector of the secondary ions at the Reflectron stage (in the text referred to as “Detector with post-acceleration”) should provide at least 9 keV of post-acceleration to the secondary ions, extracted from the sample by extracting voltages in the range of 1 to 5 keV, in order to achieve striking energy of the secondary ions of at least 10 keV under entire range of these extraction energies, aimed to achieve high detection efficiency of the heavy secondary ions with masses exceeding 1000 Da.

- The Reflectron stage should feature a connecting DN 63 CF ( NW 63 CF, 114 OD) vacuum flange, and provide additional CF vacuum flanges:

A) a flange with the size of DN 100 CF ( NW 100 CF, 150 OD) for the attachment of a small

turbomolecular pump, required for the separate pumping of the Reflectron stage of the mass spectrometer.

B) a flange with the size of  DN 40 CF (NW35 CF, 70 OD) for the attachment of the vacuum

gauge(s),

C) a flange for a high voltage feedthrough defining the potential of a Field Free Liner

- The Reflectron stage should enable straight passing of the secondary ions through the elements that generate reflecting field, without any deflection required, in order to allow the operation in linear mode, when the reflecting field is switched off and the corresponding high voltage connections earthed. At the end of the straight flight path, a flange DN 63 CF ( NW 63 CF, 114 OD) should be axially centred to allow the installation of the appropriate ion detector and consequential operation in the linear geometry.

- The Reflectron stage provides electrical vacuum feedthroughs equipped with the appropriate standard high voltage connectorts, readily available at the market,  for all the electrical supplies required to operate the Reflectron stage and the Detector with post-acceleration.

- External (air-side) electrical cables and Power supplies:

- Only the Power supply and 7 meter long cables with connectors providing low voltages to operate the preamplifier of the Detector with post-acceleration is to be included into the offer.

- High voltage cables and high voltage power supplies for the operation of the Reflectron stage and the Detector with post-acceleration will be acquired separately, they do not belong to this purchase.

OTHER REQUIREMENTS:

- the equipment should be CE compatible

- the price should include delivery charges

- documentation and manuals should be included

- technical assistance and service must be available within 7 working days after notification of a malfunction

FORM 1

5. BID  (PROFORMA INVOICE)

Details about the bidder:

| Name: |  |  |  |
| --- | --- | --- | --- |
| Address: |  |  |  |
| VAT identification number: |  |  |  |
| Bank account and the name of the bank: |  |  |  |
| Contact person: |  |  |  |
| Contact person's E-mail address: |  |  |  |
| Person responsible for signing the contract: podpis pogodbe: |  |  |  |

JOŽEF STEFAN INSTITUTE

Jamova cesta 39

1000 Ljubljana

| BID No.. Place:. Date:. |
| --- |
| On the basis of the invitation to tender No. 00386/2019 we are pleased to submit the following bid: |  |  |  |  |  |  |  |  |  |
| Type of purchasing: |  | Service |  | Material | x | Equipment |  | Construction |  |

| No. | Description of the goods | Quantity | Price | Discount % | Value |
| --- | --- | --- | --- | --- | --- |
| 1 | SUPPLY OF A REFLECTRON STAGE OF TOF MASS SPECTROMETER FOR MeV-SIMS | 1 Set |  |  |  |

|  |  |  | TOTAL |  |
| --- | --- | --- | --- | --- |
|  |  |  | DISCOUNT |  |
|  |  |  | VALUE without VAT (EUR) |  |

| Delivery time (in weeks): |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Payment terms: |  |  |  |  |  |  |  |
| Delivery term:; DAP Ljubljana (price includes delivery) the client's address). |  |  |  |
| Warranty period (at least 12 months): |  |  |  |  |  |  |  |
| Validity of the bid: |  | 31 July 2019 |  |  |  |  |  |

MANDATORY ENCLOSURE:

Brochures including a complete technical description and specification of the equipment, and the supplier’s Offer/Quotation with technical specifications and a list of components included (please, enclose it to the second part of the tender documentation)

|  |  | The bidder: |
| --- | --- | --- |
|  |  | (signature of the representative) (name and surname) |

FORM  2

4. DETAILS ABOUT THE BIDDER

Details about the bidder:

| Name of the bidder: |  |
| --- | --- |
| Address of the bidder: |  |
| Legal representative of the bidder: |  |
| Contact person: |  |
| E-mail of the contact person: |  |
| Telephone number: |  |
| Fax number: |  |
| Website: |  |
| VAT identification number: |  |
| Bank account and the name of the bank: |  |
| Person responsible for signing the contract: |  |

On the basis of a public tender »SUPPLY OF A REFLECTRON STAGE OF TOF MASS SPECTROMETER FOR MeV-SIMS« published on the Public Procurement Portal, on 18 April 2019 under publication number JN002464/2019-W01we are applying to your invitation to tender and enclose our offer documents in accordance with the Instructions for the bidders.

|  |  | The bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

The contracting authority’s note: Form must be completed and attached to the bid, if you submit bid individually.

FORM 2.1

GENERAL DATA ABOUT THE BIDDER AND THE CONSORTIUM

Details about the group leader in a joint bid:

| Name of the bidder: |  |
| --- | --- |
| Address of the bidder: |  |
| Legal representative of the bidder: |  |
| Contact person: |  |
| E-mail of the contact person: |  |
| Telephone number: |  |
| Fax number: |  |
| Website: |  |
| VAT identification number: |  |
| Bank account and the name of the bank: |  |
| Person responsible for signing the contract: |  |

On the basis of a public tender »SUPPLY OF A REFLECTRON STAGE OF TOF MASS SPECTROMETER FOR MeV-SIMS« published on the Public Procurement Portal, on 18 April 2019 under publication number JN002464/2019-W01we are applying to your invitation to tender and enclose our offer documents in accordance with the Instructions for the bidders.

Contract partners in a joint bid and description of work:

| No. | Name of the partner | Description of the work | Value | Share in % |
| --- | --- | --- | --- | --- |
| The group leader in a joint bid |  |  |  |  |
| Contract partner in a joint bid |  |  |  |  |
| Contract partner in a joint bid |  |  |  |  |
| Contract partner in a joint bid |  |  |  |  |
| Contract partner in a joint bid |  |  |  |  |

|  |  | The bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

The contracting authority’s note: Form must be completed and attached to the bid, only if you submit joint bid as

a consortium.

FORM 3

ACTING WITH OR WITHOUT SUBCONTRACTORS

Acting as a bidder, we are submitting the bid:

|  | a) individually, without subcontractors |
| --- | --- |
|  | b) with the following subcontractors: |

| No. | Subcontractors |
| --- | --- |
| 1. |  |
| 2. |  |
| 3. |  |
| 4. |  |
| 5. |  |

The bidder shall complete the table »Part of the execution of the contract, which will be performed by a subcontractor« for each subcontractor as many times as different parts of the execution of the contract will be taken by each subcontractor.

Part of the execution of the contract that will be performed by a subcontractor:

| Subcontractor |  |
| --- | --- |
| Description of work |  |
| Quantity |  |
| Value |  |
| Place |  |
| Time of completion |  |

|  |  | The bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

The contracting authority’s note: Form must be completed and attached to the bid, regardless the fact if you

submit bid individually or with a joint bid with subcontractors

FORM  3.1

THE SUBCONTRACTOR’S AUTHORISATION RELATING TO DIRECT

PAYMENTS MADE BY THE CONTRACTING AUTHORITY TO THE

SUBCONTRACTOR(S) AND CONSENT

| Name of the subcontractor: |  |
| --- | --- |
| Address of the subcontractor: |  |
| MunicipalIty of the subcontractor: |  |
| Legal representative of the subcontractor: |  |
| VAT identification number: |  |

In accordance with paragraph 5 of Article 94 of PPA-3 we declare (tick):

YES, we authorise the Contracting Authority to make direct payments;

NO, we do not require direct payments by the Contracting Authority

Subcontractors who authorize the Contracting Authority to make direct payments and above tick YES, by signing this statement give consent to the Contracting Authority to pay the subcontractors' claims against the bidder.

If the bidder will carry out the contract with subcontractors, it must provide the list of all subcontractors and type of work that the subcontractors will be undertaking, contact information and legal representatives of subcontractors and complete a Declaration of compliance with the conditions (FORM 4.1) of subcontractors in accordance with 79. Article of PPA-3 and accompanied by a request for direct payment of the subcontractor if requested by the subcontractor.

|  |  | The subcontractor: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

The contracting authority’s note: Form must be completed and attached to the bid, for each subcontractor.

The bidder, who does not act with subcontractors, does not need to submit this form.

6. FORMS FOR ESTABLISHING THE CAPACITIES OF A BIDDER

- FORM 4.1 – Declaration of compliance with the conditions for the economic

operator

- FORM 4.2 – Statement of the bidder for obtaining personal information from

Official records

- FORM 4.3 – Declaration  of compliance with the conditions for the pursuit of a

professional activity

- FORM 4.4 – Statement of technical suitability

- FORM 4.5 – List of references

- FORM 4.6 – Statement of the reference client

FORM  4.1

DECLARATION OF COMPLIANCE WITH THE CONDITIONS FOR THE ECONOMIC OPERATOR

(individual bidder, contract partner in a joint bid, subcontractor, entity whose capacity the bidder intends to use)

| Name of the bidder: |  |
| --- | --- |
| Address of the bidder: |  |
| Municipality of the bidder: |  |
| VAT identification number: |  |

We make all the statements listed below under criminal and material liability:

- The economic operator or any individual person, who is a member of its administrative,  management or supervisory body of the said economic operator, or is authorised to represent it, decide on its behalf or supervise it, has not been issued a final judgement with the elements of the following criminal offences, as indicated in the first paragraph of Article 75 of the Public Procurement Act (the Official Gazette of the Republic of Slovenia, Nos.91/2015; hereinafter PPA-3)

- We meet the obligations related to taxes and other monetary non-fiscal obligations pursuant to the law regulating of the financial administration, collected by the tax authority in line with the provisions of the country in which we have our registered office or in Member State of the Contracting Authority and we have no outstanding liabilities as at the day of submitting the bid worth EUR 50 or more. As at the day of submitting the bid, we have obtained all accounts of tax deductions for the incomes arising from employment relationships for the period of the last five years.

- On the deadline for the submission of bids we are not excluded from the public award procedures due to being included in the record of entities with negative references, as defined in Article 110 of the PPA-3.

- In the last three years prior to the expiry of the deadline for submitting the bids, we have not been fined twice for an offence related to the payment for work, about working hours, on rest, on the performance of work on the basis of civil law contracts, despite the existence of elements of an employment relationship or in connection with the employment of illegal workers with a final decision of the competent authority of the Republic of Slovenia or another Member State or a third country.

and are giving the following

CONSENT

The undersigned agrees that the Jožef Stefan Institute may, as part of the awarding procedure, obtain the data from the official records needed for checking the fulfilment of the requirements from the tender documentation relating to the »SUPPLY OF A REFLECTRON STAGE OF TOF MASS SPECTROMETER FOR MeV-SIMS«.

|  |  | The bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

FORM 4.2

STATEMENT OF THE BIDDER FOR OBTAINING

PERSONAL INFORMATION FROM OFFICIAL RECORDS

| Name of the bidder: |  |
| --- | --- |
| Address of the bidder: |  |
| Municipality of the bidder: |  |
| VAT identification number: |  |

We hereby authorise the JOŽEF STEFAN INSTITUTE, Jamova cesta 39, 1000 Ljubljana to obtain a statement from the Ministry of Justice relating to the criminal-record data on natural persons for the purpose of checking the fulfilment of the requirements of the public tender for the »SUPPLY OF A REFLECTRON STAGE OF TOF MASS SPECTROMETER FOR MeV-SIMS«, published on the Public Procurement Portal, on 18 April 2019, under publication number JN002464/2019-W01for all natural persons, authorized as legal representatives:

a)

| Name and surname: |  |
| --- | --- |
| Date of birth: |  |
| Place of birth: |  |
| Municipality of birth: |  |
| Country of birth: |  |
| Permanent/temporary address: |  |
| Signature |  |

b)

| Name and surname: |  |
| --- | --- |
| Date of birth: |  |
| Place of birth: |  |
| Municipality of birth: |  |
| Country of birth: |  |
| Permanent/temporary address: |  |
| Signature |  |

Under criminal and material accountability, we declare that the above information is true and correct. If the contracting authority has reasonable suspicion, that the bidder, irrespectively of the classification of its bid, submitted false declarations or documents, the contracting authority will submit to the National Review Commission a proposal for infringement proceedings from 112 Article of the PPA-3

This statement is an integral part of, and an annex to, the bidder’s application relating to the public tender published on the Public Procurement Portal.

|  |  | The bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

FORM 4.3

DECLARATION  OF COMPLIANCE WITH THE CONDITIONS

FOR THE PURSUIT OF A PROFESSIONAL ACTIVITY

| Name of the bidder: |  |
| --- | --- |
| Address of the bidder: |  |
| Municipality of the bidder: |  |
| VAT identification number: |  |

The bidder is for the occupation that it takes in the tender registered in one of the professional or trade registers, that are kept in the Member State in which the economic operator is established.

We make all the statements listed below under criminal and material liability:

- we are registered in one of the professional or trade registers (tick with X and complete):

the professional  registers (register name:      )

the trade registers (register name:      ),

that are kept in the Member State in which we are established.

- we have a valid registration and all the necessary licences issued by the competent authorities for performing the activity that is the subject matter hereof       OR we are member of the organisation:      .

|  |  | The bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

FORM 4.4

STATEMENT OF TECHNICAL SUITABILITY

We hereby declare that we can provide the equipment with the required technical suitability in line with all technical requirements included in the tender documentation.

Our systems enables the following TECHNICAL SPECIFICATIONS:

1. Reflectron stage of the TOF Mass Spectrometer (Short: Reflectron stage) will enable incremental

upgrade of the existing linear Time-Of-Flight (TOF) mass spectrometer for MeV-SIMS at Jožef

Stefan Institute, using existing retracting/supporting mechanics upon the simple adaptation.

- Reflectron stage will comprise of the appropriate drift space with a selectable potential (Field Free Liner), a device with reflecting electrostatic field, and detector of secondary ions at the end of the secondary ion path. The Reflectron stage is a part of high vacuum system, and its components vacuum leak-proof.

- The detector of the secondary ions at the Reflectron stage (in the text referred to as »Detector with post-acceleration«) at least 9 keV of post-acceleration to the secondary ions, extracted from the sample by extracting voltages in the range of 1 to 5 keV, in order to achieve striking energy of the secondary ions of at least 10 keV under entire range of these extraction energies, aimed to achieve high detection efficiency of the heavy secondary ions with masses exceeding 1000 Da.

- The Reflectron stage will feature a connecting DN 63 CF ( NW 63 CF, 114 OD) vacuum flange, and provide additional CF vacuum flanges:

A) a flange with the size of DN 100 CF ( NW 100 CF, 150 OD) for the attachment of a small

turbomolecular pump, required for the separate pumping of the Reflectron stage of the mass spectrometer.

B) a flange with the size of  DN 40 CF (NW35 CF, 70 OD) for the attachment of the vacuum

gauge(s),

C) a flange for a high voltage feedthrough defining the potential of a Field Free Liner

- The Reflectron stage will enable straight passing of the secondary ions through the elements that generate reflecting field, without any deflection required, in order to allow the operation in linear mode, when the reflecting field is switched off and the corresponding high voltage connections earthed. At the end of the straight flight path, a flange DN 63 CF ( NW 63 CF, 114 OD) should be axially centred to allow the installation of the appropriate ion detector and consequential operation in the linear geometry.

- The Reflectron stage will provide electrical vacuum feedthroughs equipped with the appropriate standard high voltage connectorts, readily available at the market,  for all the electrical supplies required to operate the Reflectron stage and the Detector with post-acceleration.

- External (air-side) electrical cables and Power supplies:

- Only the Power supply and 7 meter long cables with connectors providing low voltages to operate the preamplifier of the Detector with post-acceleration is  included in the offer.

- High voltage cables and high voltage power supplies for the operation of the Reflectron stage and the Detector with post-acceleration do not belong to this purchase.

OTHER REQUIREMENTS:

- the equipment will be CE compatible

- the price includes delivery charges

- documentation and manuals will be included

- technical assistance and service will be available within 7 working days after notification of a malfunction

This statement is an integral part of, and an annex to, the application to the public tender for the supply of »SUPPLY OF REFLECTRON STAGE OF TOF MASS SPECTROMETER FOR MeV-SIMS«,

|  |  | The bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

FORM 4.5

LIST OF REFERENCES

Acting as a bidder, we hereby declare that, for the past 5 years (2014, 2015, 2016, 2017, 2018), we have sold at least 3 TOF mass spectrometers with reflectron stages:

| Client, the client's contact person responsible for additional information, the relevant e-mail contact and phone number | Subject of the contract | Value of the goods (in euros without VAT) | Year (date of signing of the contract) |
| --- | --- | --- | --- |

*a minimum of 3 references

*Appendix: Confirmation of the above references with FORM 4.6 – Statement of the Reference Client

This statement is an integral part of, and an annex to, the application to the public tender for the »SUPPLY OF A REFLECTRON STAGE OF TOF MASS SPECTROMETER FOR MeV-SIMS«.

|  |  | The bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

FORM 4.6

STATEMENT OF THE REFERENCE CLIENT

Details about the reference client (confirming the reference):

| Name: |  |
| --- | --- |
| Address: |  |
| Postcode and place: |  |

Details about the bidder (requesting a reference):

| Name: |  |
| --- | --- |
| Address: |  |
| Postcode and place: |  |

We hereby state that the bidder, requesting to obtain a reference, efficiently, with respect to quality and amount, supplied us with the following goods:                           (type of goods supplied), on the basis of the contract/purchase order made with us as a client or as of       (date of signing of the copntract)

If the Jožef Stefan Institute, the contracting authority, wishes to verify our statement, they may do so by contacting                 (name and surname) at the telephone number            and/or e-mail address                .

|  |  | Name and surname of the responsible person of the reference client: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative of the reference client) |

The contracting authority’s note: The form should be photocopied for the required number of statements.

The contracting authority will consider only reference statements that will be fully completed and signed by the reference subscriber for which the provider has carried out the supply of goods, and which will also be listed in the reference list on FORM 4.5

FORM 5

7. SAMPLE CONTRACT

Jožef Stefan institute,

Jamova cesta 39,

1000 Ljubljana,

Slovenia

VAT-ID: SI55560822

represented by the Director, prof. dr. Jadran Lenarčič, henceforth referred to as “the Client”,

and

,(company name)

, (address)

,

VAT-ID:

represented by the director        (henceforth referred to as “the Supplier”)

have agreed on the following

C O N T R A C T

Article 1

SUBJECT OF THE CONTRACT

The contracting parties have established that, in line with the Public Procurement Act, the Client awarded the contract as the most cost-effective bidder with regard to the realisation of the supply requirement expressed in the public tender published on the Public Procurement Portal of the Republic of Slovenia under the reference number JN002464/2019-W01.

With this contract the Client and the Supplier agree on their mutual rights and obligations, necessary for the smooth operation of purchasing and selling the following device:

SUPPLY OF A REFLECTRON STAGE OF TOF MASS SPECTROMETER FOR MeV-SIMS

specified, in detail, in the Quotation/pro-forma invoice with the ref. number ………..... as of ………….... that is an integral part of this contract.

Article 2

VALUE OF THE CONTRACT

The value of the above equipment is …….………….. . The price does not include the VAT.

Article 3

DELIVERY TERMS

The delivery period for the equipment that is the subject of this contract is ………......... after the date of the contract.

The Supplier shall provide, without additional cost to the Client, transportation including transport insurance of the equipment to the Client’s address. The Supplier shall notify the Client of the intended delivery by fax, e-mail, in writing or by telephone, at least 3 working days prior to the delivery.

The Client shall sign the handover note after establishing that the correctly delivered equipment has been properly installed and that the Supplier presented the warranty certificates to the Client.

The Supplier shall provide the delivery of the equipment with its own professionally qualified staff or with the subcontractors, listed in tender documentation. If the Client finds that the work is carried out by the subcontractor not listed in the tender, he is entitled to cancel the contract. The Supplier reserves the right to verify any time subcontractor employees performing work. All employees are required to give credible information.

Article 4

PAYMENT TERMS

The Supplier shall issue the invoices to the Client in the following order:

- 30% payment in advance after receipt of the order and signed contract

- 70% payment payable on receipt of Invoice, at the time of delivery and technical acceptance

The Client is obliged to pay an invoice for the goods, i.e., the equipment, within 30 days of the correct issuing of the invoice.

The contracting parties agree that, in the case of a late payment, the Client has to pay the default interest.

Article 5

ACCEPTANCE

The acceptance shall be performed on the basis of the Final Acceptance Certificate (APPENDIX 1), signed by the authorised representatives of both contracting parties, upon the proper delivery of the equipment of adequate quality and quantity.

The Supplier will submit to the Client the following documentation:

- Invoice

- Delivery note (with an exact list of delivered parts, data about the gross and net weight)

- Technical documentation and instructions for use

- Precise installation instructions

- Handover note – final acceptance certificate

- Warranty certificate

The documentation must be in Slovene or English.

Article 6

SUPPLIER GUARANTEES

The responsibility for keeping the equipment undamaged is transferred from the Supplier to the Client once the equipment is delivered to the Client.

The Supplier guarantees to the Client:

- that the equipment is new, not used and free of any actual and legal defects;

- that the supplied equipment functions perfectly and that it complies fully with all the technical descriptions, characteristics and specifications, provided by the Client;

- that the Client shall acquire all the rights related to the equipment, while the Supplier shall duly perform all the obligations related to the equipment.

The Supplier provides the Client with a warrantee to cover the quality of the equipment and its smooth operation at least to an extent equal to the extent of the warrantee given by the producer of the equipment to the Supplier, i.e., ……...years following the date of the receipt of the equipment.

During the warrantee period the Supplier shall provide for emergency services responding to the Client by E-mail or telephone within 7 working day after the receipt of a notification. If an on-site repair service is needed, the Supplier shall provide it typically within 30 working days after the problem has been identified.

Article 7

ELIMINATION OF ERRORS AND SUPPLY OF SPARE PARTS

The Supplier undertakes to eliminate the errors of the delivered goods within the warranty period providing uninterrupted service to its own costs, including transport costs to a specific location. The period of notice is considered to be the time when the message reaches the supplier at the phone number or e-mail stated in this agreement, provided that it has been awarded by the Client and shall contain at least the necessary data for the identification of the goods.

The Supplier commits to provide the spare parts for at least 10 years. In the case of failure to fulfil obligations, the supplier must reimburse any additional costs and damage to the Client, which would consequently suffer.

Article 8

MAINTENANCE AND SERVICE OF THE EQUIPMENT

The maintenance of the equipment is carried out regularly, according to a predetermined schedule in line with the manufacturer's instructions and recommendations to ensure safe operation, to reduce the possibility of failures and to ensure the continuous optimal operation of the equipment. It is performed during the warranty period and after the expiration of the warranty. The price includes material, work, daily allowances and travelling costs.

Service of the equipment is a repair service that is performed in the case of equipment failure on the basis of the customer's order after the expiration of the warranty period. Service services are intended to establish such condition of the equipment in which the equipment performs its intended function. Service maintenance costs are calculated according to the actual time and material used. The Supplier informs the Client about at least the approximate costs before the service is performed.

The responsibility for keeping the equipment undamaged is transferred from the Supplier to the Client once the equipment is delivered to the Client.

The Supplier shall provide for emergency services responding to the Client within two working days after the receipt of a notification.

Article 9

ANTI-CORRUPTION CLAUSE

If it is found that, during the execution of the public tendering, on the basis of which this contract has been made, or during the execution of this contract, anyone acting in the name of – or on the account of – the Supplier, has promised, offered or given any undue advantage to the Client (the contracting authority), or those employees that had a decisive influence on the choice of the Supplier, in order:

- to win this public contract,

- to conclude this contract under more favourable conditions,

- to omit due diligence in the supervision of the contract implementation,

- to carry out any other act or omission causing damage to the Client or allowing the Client or its employees to obtain an undue advantage,

this contract will be void.

In the case of an alleged existence of the conditions from the first paragraph of this article or from a notification of the Commission for the Prevention of Corruption, or any other responsible body, relating to their alleged existence, the Client shall initiate the proceedings for invalidity of the contract from the first paragraph of this article, or any other measures, in line with the regulations of the Republic of Slovenia.

Article 10

EARLY TERMINATION OF THE CONTRACT

The Client may withdraw from the contract without obligation towards the Supplier, if the Supplier:

- does not comply with all the requirements of the Client,

- increases the prices during the validity of the contract,

- does not carry out the subject of the contract with the agreed quality or within the agreed deadlines,

- does not fulfil all of its obligations under the contract,

- in other cases and in the extent specified in this contract

and informs him with a written notice by registered mail.

The Supplier has the right to withdraw from this contract in the event of the violation of the terms of the contract by the contracting authority. In this case, the contract will expire when the contracting authority receives a written notice of withdrawal, stating the reason for the assignment by the registered mail.

In the event of withdrawal, the parties shall be obliged to fulfil the obligation so far made as agreed prior to the cancellation.

Article 11

TERMS FOR TERMINATION OF THE CONTRACT

This contract is concluded under the termination terms that will be met if one of the following circumstances is fulfilled:

- if the Contracting Authority is informed that a competent court has established with a final decision that the Suppler or its Subcontractor has infringed applicable obligations in the area of environmental, social and labour law, or

- if the Contracting Authority is informed that the competent state authority established that the Supplier or its Subcontractor has at the time of the contract infringed at least two infringements in relation to:

- payment for work,

- working hours,

- on rest,

- the performance of work on the basis of civil law contracts, despite the existence of elements of an employment relationship or in connection with the employment of illegal workers

and for which an offence has been imposed by a final decision or several final decisions of

the competent authority,

and provided that the acquaintance with the infringement and the expiration of the validity of the contract has remained for at least 6 months or if the Supplier acts with the Subcontractor, even if the Supplier does not replace or replace this subcontractor due to the established violation with the Subcontractor, in a manner determined in accordance with Article 94 of the PPA-3, and determined these contracts within 30 days from the knowledge of the violation.

If the circumstances and conditions referred to in the preceding paragraph are fulfilled, the contract shall be terminated from the date of conclusion of a new contract on the execution of a public contract for the contract in question. The contracting authority shall inform the Supplier of the date of the conclusion of the new contract.

If the Contracting Authority does not initiate a new public procurement procedure within 30 days of becoming aware of the infringement, the contract shall be deemed to be terminated on the thirtieth day after the acquaintance with the infringement.

Article 12

FINAL PROVISIONS

The contracting parties shall determine all the amendments to this contract, after a previously reached agreement, with the annexes to this contract.

The contracting parties shall resolve any disagreement relating to the realisation of this contract by common consent.

In the case that the contracting parties cannot settle a dispute by common consent, they shall abide by the decision of the competent court in Ljubljana.

This contract comes into force on the day when it is signed by both contracting parties,

The contract is drawn up in four identical copies so that each contracting party receives two copies.

| CLIENT: |  | SUPPLIER |
| --- | --- | --- |
| Jožef Stefan Institute |  |  |
| Director: |  | Director: |
| Prof. Dr. Jadran Lenarčič |  |  |
| Date: |  | Date: |

We herewith declare that the format of this sample contract is acceptable for the bidder.

|  |  | The bidder: |
| --- | --- | --- |
| (place, date) | (stamp) | (signature of the representative) |

Appendix 1: Acceptance certificate for

»SUPPLY OF A REFLECTRON STAGE OF TOF MASS SPECTROMETER FOR MeV-SIMS«

| Name: |  |
| --- | --- |
| Address: |  |
| VAT identification number: |  |
| Contact person's: |  |
| Contact person's E-mail address: |  |
| Person responsible for signing the contract: |  |
| Contract No.: |  |

| Following Delivery note Nr. |  | of |  | herewith we issue : |
| --- | --- | --- | --- | --- |

| FINAL ACCEPTANCE CERTIFICATE |  | * No. |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | place: |  |  |  |  |
|  | date: |  |  |  |  |

| Type of purchasing: |  |  |  | Service |  |  |  | Material | X | Equipment |  |  |  |  | Construction |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Delivery method: | x |  | Supplier |  |  |  | JSI arrangement: |  |  |  |  | Other: |  |  |  |  |

| Nr. | Findings |  | YES |  | NO |  |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | All the ordered parts or services have been delivered. |  |  |  |  |  |
| 2 | The quality of the supplied parts or services complies with all the terms of the purchase agreement. |  |  |  |  |  |
| 3 | The required quantity of individual components or services has been supplied. |  |  |  |  |  |
| 4 | All the required or necessary documentation is included. |  |  |  |  |  |
| 5 | Allthe manuals including the instruction for use are included. |  |  |  |  |  |
| 6 | All the other required acceptance tests have been carried out (QA) |  |  |  |  |  |
| According to the above findings, we confirm acceptance |  |  |  |  |  |  |

Additional comments:

Supplier shall receive a copy of the signed final acceptance certificate.

| The Supplier: |  | Institut “Jožef Stefan” |  |
| --- | --- | --- | --- |
| (name and surname of the supplie’s representativer) |  | (name and surname of the responsible for the equipment) |  |